## Slide 1
MEDICAL LASER SYSTEMS
Assist Prof. Dr. Lutfi Ghulam Awazli
Institute of laser for postgraduate studies
University of Baghdad

## Slide 2
TYPES OR CLASSIFICATION OF LASER

## Slide 3
TYPES OR CLASSIFICATION OF LASER
Many classifications:
1- ACCORDING TO ACTIVE MEDIUM
( Gas , Solid, Liquid, or a Semi-conductor )
2- ACCORDING TO POWER
( Low power, Intermediate power , High power )
3- ACCORDING TO WAVELENGTH
( Ultraviolet, Visible, or Infrared )
4- ACCORDING TO OUTPUT (NATURE OF EMISSION ) ( Continuous wave ,Pulsed wave laser )
5- ACCORDING TO LASER TECHNIQUE
( Contact and non-contact )

## Slide 4
Gas Laser :
Solid state laser :
Liquid dye laser:
Semiconductor Lasers:
He-Ne laser, CO2 laser
Ruby laser, Nd: YAG Laser
Pulsed Dye lasers (PDL)
Diode Lasers
1- ACCORDING TO ACTIVE MEDIUM

## Slide 5
CLASS I LASER : Low power lasers usually less than 0.4 mW.
Examples : laser printers, CD players, and diode lasers
CLASS II LASER : Low power lasers { 0.4 -1 mW }.
Examples : laser pointers, range finders
CLASS IIIa LASER : Intermediate power lasers { 1- 5 mW }
Example : laser scanner.
CLASS IIIb LASER : Intermediate power lasers either
CW wave laser power of 5- 500 mW} or pulsed (energy density of less than 10 J /cm²)
Example : dye lasers
CLASS IV LASER : High power lasers include:
(medical laser) CW lasers power of more than 500 mW or
pulsed type lasers with energy density of more than 10 J /cm² . Examples : argon laser and Nd:YAG laser,CO2 laser, …….
2- ACCORDING TO POWER

## Slide 6
Ultraviolet
Argon fluoride (Excimer-UV) Krypton chloride (Excimer-UV) Krypton fluoride (Excimer-UV) Xenon chloride (Excimer-UV) Xenon fluoride (Excimer-UV) Helium cadmium (UV) Nitrogen (UV)
visible
Helium cadmium (violet) Krypton (blue) Argon (blue) Copper vapor (green) Argon (green) Krypton (green) Frequency doubled Nd: YAG (green) Helium neon (green) Krypton (yellow) Copper vapor (yellow)
Helium neon (yellow) Helium neon (orange) Gold vapor (red) Helium neon (red) Krypton (red) Rohodamine 6G dye (tunable) Ruby (CrAlO3) (red)
Infrared
Gallium arsenide (diode-NIR) Nd:YAG (NIR) Helium neon (NIR) Erbium (NIR) Holmium: YAG (NIR) Hydrogen fluoride (NIR) Carbon dioxide (FIR) Carbon dioxide (FIR)
594 610 627 633 647 570-650 694
Wavelength (nm)
Laser Type
3- ACCORDING TO WAVELENGTH
Wavelength (nm)
Laser Type
Key:
UV = ultraviolet (200-400 nm) VIS = visible (400-700 nm) NIR = near infrared (700-1400 nm)
193 222 248 308 351 325 337
441 476 488 510 514 528 532 543 568 570
810-980 1064 1150 1504 2100 2700 9600 10600

## Slide 7
LASER BEAM OUTPUT
Lasers operate in two modes:-
1- Continuous wave mode: (CW mode)
Helium-Neon, Argon, Krypton or CO2 lasers.
2- Pulsed mode : Ruby and Nd:YAG lasers.
Pulsed laser mode may be further subdivided
according to pulse length and methods for
producing such pulses.
- Normal pulsed mode.
- Q-switched mode.
- Mode locked mode.
Power( Watts)
Time
Energy (Joules)
Time
4- ACCORDING TO OUTPUT (NATURE OF EMISSION )

## Slide 8
Lasers may be divided into two broad groups:
Continuous wave (CW) mode:
Is one whose power out put undergoes little or no fluctuation with time. It exhibits a steady flow of coherent energy. Typical examples: Helium neon, argon gas and CO2 lasers
The out put of CW lasers is measured as power in watts, referring to the rate at which work is performed, or the energy applied per unit time.
Pulsed mode:
The out put beams undergo marked fluctuations with time.
Nd: YAG and CO2 laser often, but not always, operate in the pulsed mode.
The out put of pulsed lasers is expressed in energy measured in joules. The peak power can then be calculated, equal to output energy/ pulse duration.
Pulsed laser operation may be further subdivided according to pulse length and methods for producing such pulses. The following are the basic operating modes for pulsed lasers: Normal pulsed mode.
Q-switched mode.
Mode locked mode.
LASER BEAM OUTPUT MODALITIES :

## Slide 9
Laser Output
Continuous Output (CW)
Pulsed Output (P)
Power (Watts)
Time
Energy (Joules)
Time
The output undergoes little or no fluctuation with time. It exhibits a steady flow of coherent energy. . cutting
The out put beams undergo marked fluctuations with time.
ablation + fragmentation
Examples: He-Ne , argon and CO2 lasers.
Examples: Ruby and Nd:YAG laser
The out put of pulsed lasers is measured as energy in joules.
The out put of CW lasers is measured as power in watts (1 watt =1 joule per second),

## Slide 10
Laser techniques:-
In general, there are two types of laser techniques, contact and non-contact.
1- Contact technique: There is direct contact between the fiber optic tip and the target tissue.
The tissue absorbs the radiant energy and becomes hot, resulting in the vaporization of the target cells. e.g: Nd:YAG laser
2- Non-contact technique: do not directly touch the tissue. Instead, the laser light transfers radiant energy to the tissue. Heat results when the tissue absorbs the radiant energy and the molecules in the tissue begin to move. e.g: CO2 laser
5- ACCORDING TO LASER TECHNIQUE

## Slide 11
Surface
Interstitial

## Slide 12
Invented in 1964 by Chandra K. N. Patel at Bell Labs.
Specifications:-
Type: Gas laser
Wavelength: Far infra-red (10600 nm).
Output: Generally available as continuous wave
(s.t as pulsed mode) in noncontact mode.
Delivery system: seven articulated arms
Medical applications:
- it cuts and coagulates the tissue at the same time.
It is called ( A laser scalpel ) because it cut the tissue by the high energy of laser light with very little or no effects on surrounding tissue.
- A laser beam also ablates or vaporizes the soft tissue with high water content. It ablate thin layers, one after the other, with no damage to underlying structures.
CO2 LASER
Carbon-dioxide(CO2) laser

## Slide 13
CO2 LASER

## Slide 14
CO2 laser microscopic surgery
Ultra pulse CO2 laser microsurgery system is made up of:-
Super pulse CO2 laser instrument,
Binocular operating microscope and Micromanipulator.
Can be used in delicate surgery:
- Neurosurgery,
- E.N.T,
- Ophthalmology ,
- Gynecology,
- Vascular surgery. ……………..

## Slide 15
CO2 LASER

## Slide 16
CO2 LASER

## Slide 17
CO2 LASER

## Slide 18
CO2 LASER

## Slide 19
Medical applications:
It is used in surgery for:
• Hoemeostasis
• Tumor ablation
• Vaporization of pathologic tissue
Nd:YAG Laser
Nd:YAG Laser
Neodymium : Yttrium Aluminum Garnet laser
Invented in 1961 by Leo F. Johnson at Bell Labs
Specifications:-
Type: A solid state laser.
wavelength :1064 nm (Near infra-red).
Output: Available as CW or pulsed in contact mode
Delivery system: Fiber optic
- It is used in open and endoscopic surgery
Fiber optic

## Slide 20
Nd:YAG Laser

## Slide 21
Nd:YAG Laser

## Slide 22
Nd:YAG Laser

## Slide 23
Nd:YAG Laser

## Slide 24
Diode Laser
Diode Laser
Invented in 1962 by Robert Hall at General Electric.
Type: Semiconductor Lasers
Wavelengths: Infra-red (810nm, 940nm, 980nm)
Operating modes: continuous (CW) and pulsed. . In contact or non-contact mode.
Delivery system: Fiber optic
Medical applications:
its excellent absorption in hemoglobin, melanin and water.
it can be used for open as well as endoscopic or endoluminal applications

## Slide 25
Diode Laser( diomax)
ENT Surgery
• Vascular anomalies
• Nasal polyps/tumors
• Nasal concha vaporization
• Tonsillectomy
Maxillofacial Surgery
• Mucosal diseases
• Oral tumors
Dermatology/Vascular surgery
• Vascular anomalies
• Hemangiomas
• Nevi
• Tumors
• Warts
General Surgery
• Endoscopy
• Laparoscopy
• Gynecology
Endovenous Laser Treatment (EVLT)
• Great saphenous vein
• Small saphenous vein
• Branch varices

## Slide 26
Diomed Delta Laser
Dornier D940 Laser
Endovenous Laser Treatment (EVLT)

## Slide 27
Diode Laser in dental surgery

## Slide 28
(Ho:YAG) solid state, pulsed | Laser type
2100 nm | wavelength
350 μsec | Pulse duration
0.2 -3.5 J | Energy Range per pulse
80 watt | Maximum power
Red diode laser 635 nm, continuous maximum power (5 mW) | Aiming Beam
Single pulse mode (5-60 HZ) and Double pulse mode (3-30 HZ) | Lasing mode
Trimedyne holmium optical fiber | Laser Delivery
Self-contained water to air heat exchange to ambient air | Cooling
303 Kg (670 Ib) | Approximate weight
Holmium (Ho):YAG laser

## Slide 29
The four adjustable laser parameters include:-
1. Pulse Mode (single /double)
2. Pulse Repetition Rate (Hertz)
Single pulse mode ( 5- 60 Hz)
Double pulse mode (3- 30 Hz)
3. Power (Watts)
Maximum power 80 watt
4. Set Time (CL or seconds)
1
2
3
4
Ho:YAG laser

## Slide 30
Orthopedic/Arthroscopic Applications:
Intradiscal Decompression
Thermal Annuloplasty
Fatty tissue removal
Discectomy
Foraminoplasty
…………
Medical applications of Holmium:Yag laser :
Urological Applications:
Ureteral Stone
Badder Stones
Kidney stones
Bladder Tumor
Ureteral Tumor
Condyloma Acuminata
Ureteral Stricture
Bladder Neck Contracture

## Slide 31
Parameters of Laser Radiation:
Wavelength (nm) : Absorption + Depth of Penetration
Power( P ): The rate at which the energy is delivered (W)
Exposure time (t) :
Pulse duration, (s)
Energy( E )= P x t
Spot size (cm).
Power Density (Irradiance)=P/A (W/cm²)
Energy density (Fluence)= E/A (j /cm²)
PRR is the pulse repetition rate,( i.e: number of pulses per second )